Gangarosa Department of Environmental Health Science Laboratory

Standard Operating Procedure

Automated Sample Prep-LCHRMS

Date: 12/01/2023

Revisions: 12

Prepared by: Catherine Mullins, Maria Cardelino

Reviewed by:

Douglas Walker, Associate Professor                     Date:

Revision History

SOP Details and Version Information:

| Created By: | Catherine Mullins, Maria Cardelino | Date: 12/01/2025 |
| --- | --- | --- |
| Reviewed By: | Douglas Walker, Associate Professor | Date:12/04/2025 |
| Approved By: | Douglas Walker, Associate Professor | Date: 12/04/2025 |

Table of Contents

Overview

Sample Aliquoting

LC Supernatant Transfer

Trouble Shooting

Overview

Purpose of SOP: This procedure details steps for preparing biofluids such as plasma, serum, urine, and follicular fluids for LC analysis using Opentrons Flex automated liquid handlers.

Required BioRAFT Safety Training:

- BBP for Research

- Biosafety

- Research Lab Safety

Cautions

- Sample preparation involves hazardous chemicals and biohazardous materials. Technician must be familiar with the SDS forms for all chemicals and must follow proper procedures for handling biohazardous materials.

- PPE (eye protection, lab coat, and gloves) should be worn throughout processing the samples to prevent contamination of samples and exposure to the technician. Gloves should be changed when they become contaminated or when touching items such as cell phones not directly involved in sample prep. If computer or other items do become contaminated, clean with 70% ethanol.

- Clean work surfaces daily with 70% ethanol.

- Waste materials should be collected and disposed of in biohazardous waste bins.

Definitions/ explanations

- Water blanks: These are prepared and processed exactly like samples, except using an identical volume of LCMS grade water in place of biofluid.

Abbreviations

- DWP: Deep well plate.

- LC: Liquid chromatography

- QA/QC: Quality Assurance/ Quality Control. Plasma or urine pools 1 (young) and 2 (older) are typically used, but different biological sample types may require the use of different types of QAQCs that more closely match the study sample type.

Equipment

| Protocol Name | Manufacturer | Item |
| --- | --- | --- |
| Opentrons | Opentrons | Flex liquid handling robot |
| Pipette | Opentrons | 1000ul 96-channgel pipette |
| Centrifuge | Thermo | *Thermo |
| ThermoMixerC or Vortex | Eppendorf | ThermoMixerC |
| Station One | JDM Lab Solutions | StationOne rack thawing station |

Important Notes:

- Samples prepared for LC analysis must be kept cold. Use chilled aluminum blocks when aliquoting from vials, keep thawed samples in the refrigerator when not actively aliquoting, and cap and return samples to their original box in the freezer as soon as possible after aliquoting.

- All solvents used for LC sample preparation must be kept at 4˚C.

- When volatile solvents or samples extracts are involved, do not perform labware position checks using exposed solvents or sample plates. Rather, use empty or extra labware. Labware used for position checks must not have been formerly heat-sealed, as this affects the labware height.

- A sample prep worksheet must be filled out for each sample preparation run (see Sample Prep Worksheet).

- Opentrons robot/app upgrades require retesting of all protocols. LC level: runs on 8.6.0.

Preliminary Software Requirements: Each Opentrons user must perform these one-time tasks in order to operate the Opentrons from their user setting on a computer.

- Download the Opentron application

- Ensure that the version is 8.6.0 (LC). Download the .exe file from this link to select the correct software version. Make sure to update both the application and the robot!

- Install Notepad++ from the Emory Software Center (see your desktop).

- Import all custom labware definitions from the CLUES OneDrive folder into your Labware library in your Opentrons app.

  - Download all .json files

  - Open Opentrons app and select “labware” from the lefthand panel.

  - Select “Import” from the top right corner, then locate your file.

- Map the RSPH projects drive (“R-drive”) and your home drive (instructions HERE). This is where the instrument technician will create run folders for each study, which will include plate loading order files and the Opentrons protocols the sample prep technician should use for a given study.

Sample Aliquoting

Overview: This is the first step of sample preparation. The LC plate receives 30uL of sample, followed by 90ul of extraction solvent.

Materials:

| Protocol Name | Manufacturer | Item | Item number |
| --- | --- | --- | --- |
| Wellplate | Thermo Scientific | Nunc 384 wp, 250ul | 269390 |
| Heat Seal | Thermo Scientific | Thermo Scientific™ Easy Pierce Heat Sealing Foil | AB-0757 |
| Reservoir | Opentrons | 15ml Nest Reservoir, 12 Columns | 999-00078 |
| 200ul Pipette Tips | Opentrons | 200ul Filter Pipette Tips | 991-00111 |
| 50ul Pipette Tips | Opentrons | 50ul Filter Pipette Tips | 991-00110 |
| 15mL tube | Corning | Falcon 15 mL PP Conical Tube | 352097 |
| 50mL tube | Corning | Falcon 50 mL PP Conical Tube | 352098 |
| Acetonitrile | Fisher | UHPLC-MS Acetonitrile* | A956-1 |
| Water | Fisher | UHPLC-MS Water* | W8-1 |
| Methanol | Fisher | UHPLC-MS Methanol* | A458-1 |

*UHPLC grade solvents are used instead of Optima LCMS because of lower PFAS contamination.

Pre-Prep:

- Prepare solvent mixture of 60:40 (v/v) Methanol/Acetonitrile. This may be prepared in bulk and stored in a dedicated glass bottle at 4˚C. The bottle must be cleaned before adding solvent. Cleaning procedure: bottle will be rinsed 3x. Each rinse will be performed 1x in the following order:

    - 1:1 acetonitrile/basic water

    - UHPLC methanol

    - UHPLC acetonitrile

- Prepare Flex pipette tips for the week:

  - Protocol 1 and 2: Arrange 50uL tips in 10 columns (sample) and 5 columns (QAQC)

  - Protocol 1: Arrange 200uL tips in 2 columns for extraction solvent addition

  - Protocol 2: Water tips will be reused only throughout the week. The black tip holders are labeled; refill with 10 columns of fresh tips each week.

- Locate the Flex methods for your study. Instrument protocols must be saved within the study folder’s run list. Ensure you are running only protocols from the study folder. Path: R:\diwalke\LC\Run_Lists

- Thaw Samples: 30min-3hrs (depends on sample volume and matrix), StationOne in cold room, 4°C. While samples are thawing:

  - Start filling out LC Sample Prep Worksheet. Use this throughout the procedure to track timing and notes/observations.

  - Turn everything on:

    - Opentrons and computer

    - Set temperature modules, centrifuge, and thermomixers to 4°C

    - Preheat heat-sealer to 165°C.

  - Upload Opentrons protocols from the study’s Run Lists folder.

  - Set out labware and perform labware position check. Cover all labware/tips until ready to start protocol. Make sure the tips are emptied into biohazard.

  - Set out a tray of wet ice that can hold 3 matrix racks. Cover ice with foil.

- Prepare Extraction Solvent

  - Measure 24mL of 60:40 MeOH/ACN mixture into a 50ml Falcon tube using a clean 25mL graduated cylinder

  - Add 400uL of PFAS ISTD after a vortex. Located in -20˚C freezer near the Flex.

  - Add 400uL of LC ISTD after a thaw and vortex. Located in -20˚C freezer near the Flex.

  - Place used vials in “used ISTD box”

  - Vortex Falcon tube

  - Place Falcon tube in 4˚C

- Scan Plates and Create Sequence: Using Micronic scanner, scan in plate to create plate loading order document. The output location can be specified in the Micronic software.

  - Naming convention: CLU####_Batch#_Rack# or CLU###_Batch#_QAQC

  - Save to: R:\diwalke\LC\Run Lists\Plate_Scans

  - Confirm scanned in IDs against the plate-loading order in the Run Lists folder

  - Create sequence using the sequence generator app. Sequence can be created when samples are in the vortex.

- Vortex thawed samples: Thermomixer-C 4˚C/1100 RPM/10min

Ensure cover is secured over caps. Otherwise, the caps will be unscrewed!

- Centrifuge sample plates: 4˚C/4500 RPM/2min

  - Ensure both samples plates are of equal mass. Adjust mass by adding vials filled with water located near the balance.

  - Note: QAQC plate need not be centrifuged unless the QAQC material is not in the bottom of the vial. If centrifuge is required, centrifuge same as the samples and make a note in the worksheet.

  - Decap QAQC plate while samples are centrifuging

    - Place caps in clean, designated tray.

    - Replace cover on each rack after ensuring it is dry.

    - Place rack on top of the foil-covered wet ice.

- Run Protocol 1 on Opentrons (≈10min).

  - Add 12ml of extraction solvent (ES) to the first 2 columns of the 12-column Nest Reservoir. Use 15mL Falcon tube for measurements and deliveries.

  - De-cap all sample plates at once during ES addition:

    - Place caps in clean, designated trays.

    - Replace cover on each sample rack after ensuring it is dry.

    - Place racks on top of the foil-covered wet ice.

  - Protocol order:

    - Extraction solvent addition

    - *Pause*

    - Place 1st sample rack, uncover, + resume

    - *Pause*

    - Place 2nd  sample rack, uncover, + resume

    - *Pause*

    - Place QAQC rack, uncover, + resume

    - Notes:

      - All pipette tips should be set out at once.

      - As each rack of samples finishes, cover the rack and place on ice and place the next rack uncovered into position.

      - Wait until the protocol is done and sample extracts are equilibrating before attempting to recap the sample vials. It is important to keep samples at 4˚C throughout, but it is also important to WATCH THE ROBOT THE WHOLE TIME.

- Heat seal receiving plate: (1.5 seconds, 165C) - Use medium foil seals!!! Press with roller to ensure all wells are sealed! For first press: ensure P1 is in the back right, then turn around for the 2nd seal.

- Vortex: 4°C/1100RPM/10min.

During the vortex/equilibration:

  - Cap remaining samples and return to –80˚C

  - Save leftover ES in designated Falcon tube for CM.

  - Update worksheet.

  - Discard used tips and leftover QAQCs

  - Label the C18 and HILIC well plates

  - Set out labware and perform labware position check for final protocol. Cover all labware/tips until ready to start protocol. Make sure the tips are emptied into biohazard.

- Centrifuge: 4°C/4500RPM/2min. During this time, proceed to the LC Supernatant Transfer and begin following the steps in the procedure. Ensure plate is balanced with a plate of equal mass.

LC Supernatant Transfer

Materials

| Protocol Name | Manufacturer | Description | Part number |
| --- | --- | --- | --- |
| Wellplate | Thermo Scientific | Nunc 384 wp, 250ul | 269390 |
| Plate mat | Thermo Scientific | Thermo Scientific™ Easy Pierce Heat Sealing Foil | AB-0757 |
| Reservoir | Opentrons | 195ml Nest Reservoir | 999-00078 |
| 200ul Pipette Tips | Opentrons | 200ul Filter Pipette Tips | 991-00111 |
| 50ul Pipette Tips | Opentrons | 50ul Filter Pipette Tips | 991-00110 |
| Water | Fisher | UHPLC-MS Water | W8-1 |

Procedure: This assumes preparations were already completed during equilibration of the extracts.

- Using a clean 50ml Falcon tube, add 45ml of solvent to the reservoir; tap out any bubbles, if present:

  - C18: Water – reservoir in C2

- Click “Start Run.”

- The protocol will pause after the dilution solvent have been delivered for you to place the centrifuged sample plate in D1, uncover, and resume the protocol.

- After transfers are complete, heat seal receiving plate. (165C/1.5sec). Start with P1 in the back, then rotate.

- Centrifuge each plate. Note: C18 and HILIC plates will have a different mass—ensure balance plates match the masses.

- Vortex: 4°C/2000RPM/10min.

- Place sample plate in autosampler as directed by instrument technician.

- Deliver completed worksheet to instrument technician and discuss any sample prep issues which may affect analysis.